## Slide 2: Tissue-Engineered Skeletal Muscle Stimulator
Syed Asaad Hussain
Spencer Ryan Keilich
Stephanie Jo Lindow
Shreyas Renganathan
Notes: Asaad

## Slide 3: Overview
Background
  Significance
  Drug Approval Process
  Background Research
  Current Devices
Project Direction
  Problem Statement
  Goals
  Objectives and Constraints
  Design Specifications
Alternative Designs
  Design Evaluations
Final Design
  Final Design Diagrams
Proof-of-concept
  Design Validation
  Demonstration (Video)
Conclusions & Recommendations
Notes: Asaad

## Slide 4: Background

## Slide 5: Significance
Serious injuries and disease can lead to the loss or damage of skeletal muscle tissue
  Example: Muscular Dystrophy affects 32,000 people between the ages of 5-24 annually in the USA [1]
  Annual cost of treatment is $4,600/person [1]
In vivo animal models lack the essential genetic component to model disease
  Example: Facioscapulohumeral Muscular Dystrophy
There is a need for accurate human in vitro skeletal muscle models
[2]

## Slide 6: Drug Approval and Our Device
Drug Discovery
Pre-Clinical
Clinical
Our Device saves money and time
2-10 Years
3-6 Years
6-7 Years
90% Failure

## Slide 7: Muscle Formation from Cells

## Slide 8: Key Characteristics of Muscle
Anchorage
Stimulation

## Slide 9: Current Devices
[1]
[1]
[2]
[3]

## Slide 10: Problem Statement
No in vitro system that produces accurate models of in vivo skeletal tissue.
Current Systems:
Device | Sample Quantity | Cell Density (cells/construct) | Construct size | Anchorage | Standard Incubator | Stimulation Location
Vandenburgh | 8 | 1-4 × 106 | 25mm | Dog-bone | No | In incubator
Vandenburgh | 24 | 0.4 x 106 | 176.5mm^2 | 2D tissue | No | In incubator
Myomics | 96 | 0.2 x 106 | 3.5mm | Dog-bone | No | No stimulation
Powell | 6 | 1×106 | 21.4mm | Dog-bone | Yes | In incubator
Christ | 6 | 2 x 106 | 15mm | Clamped | No | Out of incubator
Dennis | 4 | 0.5 x 106 (Growth Media) | ~40mm | Sutured | Yes | Out of Incubator

## Slide 11: Flaws in Current Systems
Device | Sample Quantity | Cell Density (cells/construct) | Construct size | Anchorage | Standard Incubator | Stimulation Location
Vandenburgh | 8 | 1-4 × 106 | 25mm | Dog-bone | No | In incubator
Vandenburgh | 24 | 0.4 x 106 | 176.5mm^2 | 2D tissue | No | In incubator
Myomics | 96 | 0.2 x 106 | 3.5mm | Dog-bone | No | No stimulation
Powell | 6 | 1×106 | 21.4mm | Dog-bone | Yes | In incubator
Christ | 6 | 2 x 106 | 15mm | Clamped | No | Out of incubator
Dennis | 4 | 0.5 x 106 (Growth Media) | ~40mm | Sutured | Yes | Out of Incubator

## Slide 12: Goals
Tissue Anchorage
Tissue anchorage representative of muscle-myotendinous junction attachment
Minimizing Resources
Use of minimal functional units of engineered muscle, to minimize seeded-cell quantity and maximize throughput
Mechanical Stimulation
Stimulation in in vivo like conditions
Mechanical stimulation representative of in vivo passive and active muscle movement
Goal: Design an in vitro stimulation system that produces accurate models of in vivo skeletal muscle tissue in terms of cell density and tissue maturation by focusing on

## Slide 13: Objectives and Constraints
Objectives:
Reliability and Consistent Results
Mechanically Stimulate Tissue
Efficient
  Reusable
  High throughput
User Friendliness
Constraints:
Sterilizable
Must withstand in vitro culture conditions
  37 oC
  99% humidity
  5% CO2
Non-cytotoxic
Operate while in incubator

## Slide 14: Design Specifications
Stimulation Specs:
Mechanical Stimulation
  Uniaxial strain
  Controllable strain between -50% to 50%
  Static and cyclic capabilities
  Controllable strain rate
  Controllable frequency of stimulation
Tissue Culture Specs:
Tissue Structure
  Dog-bone shape
  Minimal Functional Unit
    Length: 3-4 mm
    Diameter: 1 mm
Cell Source
  C2C12 (Myoblast)
  500,000 cells/120uL

## Slide 15: Alternative Designs

## Slide 16: Motor Driven

## Slide 17: Fluid/Hydraulic Driven

## Slide 18: Vibration Driven
[1]

## Slide 19: Materials Design

## Slide 20: Design Evaluation
Design Means | Strain -50% to 50% (with initial length of 3.5 mm) | Representative of muscle bone attachment (uniaxial) | Expandable for high throughput | Electric/power components outside incubator | Controllable by user settings | Number of Functions Fulfilled
Stepper Motor and Cam | Yes | Yes | Yes | No | Yes | 4/5
Screw Driven | Yes | Yes | Yes | No | Yes | 4/5
Vibration | No | Yes | Yes | No | No | 2/5
Scissor Linkage | Yes | Yes | No | Yes | Yes | 4/5
Hydraulic/ Pneumatic | Yes | Yes | Yes | Yes | Yes | 5/5
Piezoelectric Material | Yes | No | Yes | Yes | Yes | 4/5
Thermo-responsive | Yes | No | Yes | Yes | No | 3/5
Conductive Polymer | yes | No | Yes | Yes | Yes | 4/5

## Slide 21: Final Design Prototype Progression

## Slide 22: Final Design

## Slide 23: Final Design: Plates
+
=
Bottom Plate
Top Plate

## Slide 24: Final Design: Well and Post Design

## Slide 25: Final Design: Syringe Pump

## Slide 26: Final Design: Complete

## Slide 27: Tissue Culture
heating block (50oC)
Flood outer wells with differentiation media

## Slide 28: Tissue Culture
30min @50oC
heating block (50oC)
NIPAAm

## Slide 29: Tissue Culture
Fibrinogen/DPBS(-)
60min @ 37oC
Fibrin gel
C2C12 cells in Thrombin + Ca2+
Myoblast cells
heating block (37oC)

## Slide 30: Tissue Culture
Construct Formation
Tissue Construct

## Slide 31: Tissue Culture
30min @ 4oC
DPBS(+) rinse (100μL)
Aspirate media
Aspirate NIPAAm

## Slide 32: Tissue Culture
DPBS(+) rinse
Flood outer wells with diff. media
Add diff. media

## Slide 33: Tissue Culture
Stimulation

## Slide 34: Tissue Culture: Flow Diagram
30min @50oC
Fibrinogen
60min @ 37oC
Overnight incubation (37oC)
heating block (50oC)
Flood outer wells with diff. media
30min @ 4oC
heating block (50oC)
NIPAAm
Aspirate media
Aspirate NIPAAm
DPBS(+) rinse
Flood outer wells with diff. media
Add diff. media
Stimulation
heating block (37oC)
Flood gently (diff. media)
Fibrin gel
C2C12 cells in Thrombin + Ca2+
Myoblast cells
Construct Formation
Tissue Construct

## Slide 35: Gold Standard Comparison
Parameter | MQP | Myomics
Sample Quantity | 96 | 96
Cell Density (cells/volume) | 5 x 105 / 120 μL (C2C12) | 2 x 105 / 120 μL (primary myoblasts)
Construct size | 3.5mm | 3.5mm
Anchorage | Dog-bone | Dog-bone
Stimulation Location | In incubator | n/a
Number of Samples Stimulated at a Time | 96 | 0

## Slide 36: Operational Cost Analysis
MQP | Estimated Cost:
Full Device | $203.89
NIPAAm | $87.10/gram
Fibrinogen (bovine) | $170.50/gram
Thrombin (human) | $288.50/150UN
Cell culture agents | $187.30
TOTAL: | $937.29
 | Estimated Cost:
96 well culture plate with PDMS molds | $350.00
NIPAAm | $87.10/gram
Fibrinogen (bovine) | $170.50/gram
Thrombin (human) | $288.50/150UN
Cell culture agents | $250.00
TOTAL: | $1,146.10

## Slide 37: Strain/Volume Relationship
Chart: Amount Strained by Volume Change
dry | 0 | 0.11714 | 0.26571 | 0.17714 | 0.045714 | -0.11429 | -0.29057 | -0.20571 | -0.031429
media | -1.31e-16 | 0.13569 | 0.27139 | 0.19174 | 0.014749 | -0.13864 | -0.24484 | -0.17994 | -0.011799

## Slide 38: Video Demonstration

## Slide 39: Proof of Concept
Device Verification

## Slide 40: Tissue Results
Mechanical strain tests
  Exercise: 5%-15%
  15 minutes
  1 cycle/minute
  4 days
Stimulated vs. Control
  H+E Stain
  Myosin Stain

## Slide 41: H+E Tissue Results
Control Tissue
Stimulated Tissue
50 µm
50 µm

## Slide 42: Myosin Tissue Results
Control Tissue
Stimulated Tissue
50 µm
50 µm

## Slide 43: Conclusions & Recommendations

## Slide 44: Conclusions
DEVICE
Produces reliable and consistent results
Uniaxial mechanical stimulation of tissue
Efficient: High Throughput and Reusable
User friendliness
Market Comparison
Higher functionality
Device: $203<Others: $350
Higher content testing
No need for live animal model
Observation of human tissue interaction
Deliverable:
A mechanical stimulator of skeletal muscle tissue that produces tissue samples with improved fiber alignment, increased fiber density, and more accurate modeling of in vivo tissue when compared to non-stimulated in vitro tissue constructs.

## Slide 45: Recommendations
Perform more stimulation experiments and statistically validate results
Adapt to engineered tissue by cell self-assembly
  Increased fiber density
  More alike native tissue
Optimize media composition
  Tissue maturation
  Hypertrophy
  Fiber size
Replace post to PDMS
Replace MED610
Use more rigid tubing to reduce error
Electrical stimulation component
Test Recommendations:
Device Recommendations:

## Slide 46: Questions?

## Slide 47: References
Page, R. (2014). Client Statement.
“Tissue engineering and regenerative medicine”, 2012. Osaka Institute of Technology, from http://www.oit.ac.jp/english/engineering/img/bioEngneer/pht_bio_02.jpg
Guilak et al. (2003). Functional Tissue Engineering. Springer-Verlag New York, Inc.
Aschettino, M., Delfosse, S., Larson, K., Quinn, C. (2011). BioMimeticSkeletal Muscle Tissue Model. Major Qualifying Project. https://www.wpi.edu/Pubs/E-project/Available/E-project-042512-122709/unrestricted/RLP-1101_Biomimetic_Skeletal_Muscle_Tissue_Model.pdf
Milano, Shaun. (2014). Allegro MicroSystems, LLC.
Vandenburgh, H., Tatto, M. D., Shansky, J., Lemaire, J., Chang, A., Payumo, F., ... & Raven, L. (1996). Tissue-engineered skeletal muscle organoids for reversible gene therapy. Human gene therapy, 7(17), 2195-2200. http://online.liebertpub.com/doi/pdf/10.1089/hum.1996.7.17-2195
Vandenburgh, H. H., Hatfaludy, S., Karlisch, P., & Shansky, J. (1991). Mechanically induced alterations in cultured skeletal muscle growth. Journal of biomechanics, 24, 91-99.
Powell, C. A., Smiley, B. L., Mills, J., & Vandenburgh, H. H. (2002). Mechanical stimulation improves tissue-engineered human skeletal muscle.American Journal of Physiology-Cell Physiology, 283(5), C1557-C1565.
Radisic, M. (2005). U.S. Patent No. 20050112759. Washington, DC: U.S.
Dennis, R. (2001). U.S. Patent No. 6114164. Washington, DC: U.S.
Christ, G. (2013). U.S. Patent No. 20130197640. Washington, DC: U.S.

## Slide 48: Appendices

## Slide 49: Appendix: Cytotoxicity
Materials:
  Polystyrene (Control), Polypropylene (Blocks), MED610 (3D Printed), and Polyethylene (tubing)

## Slide 50: Appendix: Cytotoxicity Graphed
Chart: Percentage of Cells Alive after 7 Days of Media Exposure
| Polystyrene (control) | Polystyrene (control 2) | Polypropylene (1 Day incubated) | Polypropylene (2 Days Incubated) | Polypropylene (4 Days Incubated) | Polypropylene (6 Days Incubated) | MED610 (1 Day incubated) | MED610 (2 Days incubated) | MED610 (4 Days incubated) | MED610 (6 Days incubated)
 | 81.3559 | 82.963 | 88.3721 | 81.8966 | 88.8889 | 82.3529 | 81.6092 | 82.1739 | 74.5098 | 61.1111
Experimental groups compared via ANOVA analysis p-value: 0.092

## Slide 51: Appendix: Myogenic Potential
C2C12 Cell Differentiation
Phase contrast
Day 9
Fluorescence
  Myosin & DAPI

## Slide 52: Appendix: Tissue Results
Control Tissue 1
Control Tissue 2
Experimental Tissue 3
Experimental Tissue 4
Experimental Tissue 4
Experimental Tissue 1

## Slide 53: Appendix: Tissue results analysis
Control Tissue 1
Control Tissue 2
Experimental Tissue 3
Experimental Tissue 4
Experimental Tissue 4
Experimental Tissue 1

## Slide 54: Appendix: Thermal Expansion Testing

## Slide 55: Appendix: Syringe Pump Relations

## Slide 56: Appendix: Syringe Pump Relations
Calculations: Desired Strain:
Input: Desired strain (ε)
Output: ∆Vo required to reach desired strain
(1)
Volume of Cylinder:
(2)
(3)
(4)

## Slide 57: Appendix: Device Verification Test
Dry Test
Media Test

## Slide 58: Appendix: Drug Approval Process
Notes: Highlight Phase II  this device will give more accurate data at phase II
90% of drugs fail in phase to clinical trials (we are bridging this gap)

## Slide 59: Appendix: Marketability
Stakeholders
Researchers and Pharmaceutical Companies
Cheaper and more efficient means of testing
No live animal or human testing required
Patients with Muscular Diseases
Improved treatment methods for diseases
Observation of human tissue interactions
Impact
Notes: Stakeholders:
  Researchers and pharmaceutical companies searching for the treatment of muscular defects
  Patients with muscular disease
How does it impact them?
  Cheaper and more efficient means of high throughput testing
  No live animal or human testing required
  Observe human tissue interactions

## Slide 60: Appendix: Marketability
Engineering and Biological Applications
Improve discovery rate
Improve treatment of muscle defects
Entrepreneurial Applications
Reduce cost
Easily manufactural and distributable
Produce more accurate representations of ANY muscle tissue
Open market and field to smaller organizations
Open product to non-consumers

## Slide 61: Appendix: Objective Tree
Stimulator
Reliability
Consistent Results
High Throughput
Minimal Functional Unit
User Friendly
Automated
Programmable
Reusable
Function
Stimulation
Accuracy
Mechanical
Dogbone
Uniaxial
Efficiency
Cyclic and Static
Culture in Device

## Slide 62: Appendix: Device Pictures

## Slide 63: Appendix: Laboratory Protocols
Diagram from main presentation
Page Lab cell protocols
Strain regimen
1-2% strain at 1 cycle per minute for 15 minutes on stim day 1
2-4% strain at 1 cycle per 2 minutes for 15 minutes on stim day 2
2-4% strain at 1 cycle per minute for 15 minutes on stim day 3

## Slide 64: Appendix: Degenerative Muscle Disease
Diseases:
  Muscular Dystrophy
    Facioscapulohumeral
    Duchenne MD
  Multiple Sclerosis
  Amyotrophic Lateral Sclerosis
  Myasthenia Gravis
Common Symptoms:
  Loss of muscle mass
  Loss of coordination
  Signal transduction problems
  No cure
Average cost of care (MD):
  $18,930 per 4 years (per case)
Emery, Alan E. 1994 M. Muscular Dystrophy
"Muscular Dystrophy: Hope Through Research," 2013 NINDS. No. 13-77

## Slide 65: Appendix: Key Characteristics of Muscular Tissue
[1] Brock, R. “NIH study uncovers details of early stages in muscle formation and regeneration. May 2013.
[1]
Stimulation
Fiber Alignment
Anchorage
Muscle-Tendon Attachment

## Slide 66: Appendix: Tissue Physiology

## Slide 67: Appendix: Constraints (incubator conditions)
Device needs to withstand incubation
Muscle cells passively contract and rupture
Cells culture complications
  Protein denaturation
    Limits temperature
    pH
    cell density
  Viability range
    Temperature
    pH
    Density

## Slide 68: Appendix H: Raw data from tests