## Slide 1
Takes baseline pressure readings from the strain gauges
Translates strain gauge value to degree position for servo motor to establish its baseline position
Measures the patient’s average heart rate at a given time with the pulse oximeter
Based on the measured value, the servo motor turns. If the heart rate is:
  Less than 65 BPM, the servo turns clockwise 5 degrees, tightening the cast
  More than 110 BPM , the servo turns counterclockwise 5 degrees, tightening the cast
When the button is pressed, another reading from the strain gauges and translation to the degree position is taken and compared to the baseline
Based on the variation from the baseline, the servo motor will rotate to return the cast back to the baseline value, and the servo’s new degree position is recorded
LCD displays the current heart rate of the patient, along with the new servo position if changed above
Repeat the code from step 3
Sensor and Servo Enhanced Smart Arm Cast
Aedan Fahy (ME), Achilles Gikas (ME/BME), Shawn Mills (ME), Shaniqua Miller (ME)
Advisor: Mehul Bhatia (MME) Co-Advisor: Songbai Ji (BME)
OBJECTIVES
MOTIVATION
ACKNOWLEDGEMENTS
I: Ideate a potential product that serves as a smart wearable device for consumers
II: Design a model with the use of mechanical equipment and engineering principles
III: Develop the model to create a functional prototype of the smart wearable cast
IV: Test, analyze results, redesign, and redevelop the product for improvement
2.6 million emergency room visits per year due to fractures [1]
Current casting processes use skin-irritating and non-reusable material
Cast applications require tedious multi-step methods [2]
The healthcare industry lacks smart devices accessible to patients
Sensor-based technology will efficiently and effectively fixate an injured wrist autonomously [3]
FUTURE WORK AND CONCLUSIONS
REFERENCES
[1] N. N. O’Hara, D. S. Kringos, G. P. Slobogean, Y. Degani, and N. S. Klazinga, “Patients Place More of an Emphasis on Physical Recovery Than Return to Work or Financial Recovery,” Clin Orthop Relat Res, vol. 479, no. 6, pp. 1333–1343, Jun. 2021.
[2] The Center Oregon. (2022, November 11). What to expect when you get a cast. The Center - Orthopedic & Neurosurgical Care & Research. https://www.thecenteroregon.com/medical-blog/joint-injury-cast-care/
[3] R. Huggenberger and M. Detmar, “The Cutaneous Vascular System in Chronic Skin Inflammation,” J Investig Dermatol Symp Proc, vol. 15, no. 1, pp. 24–32, Dec. 2011.
TESTING AND RESULTS
Our project team conducted work throughout the academic year alongside our Mechanical Engineering advisor, Professor Mehul Bhatia Ph.D. Our additional project co-advisor from the Biomedical Engineering Department is Professor Songbai Ji Ph.D. We are also grateful to conduct work using resources from the MQP Laboratory with the help of James Van Mulligan.
CAST MATERIAL AND DESIGN MODELING
Doctor applying short arm cast to a patient
CAST SENSORS AND EQUIPMENT
Testing after assembly and wiring included:
[Pre-Code] Calibration Strain Gauges, Servo Motor, and Pulse Oximeter to make sure wiring was successful
[Code Part 1] Baseline Test on Patient's Arm to maintain pressure overtime (1 hour per test)
[Code Part 2] Testing of Pulse Oximeter to detect patient’s BPM and activation of servo motor to adjust tightness if heart rate is not in the 65-110 BPM range
[Code Part 3] Testing of Baseline Pressure deviation detection from Strain Gauge and adjustment from servo motor
ARDUINO WIRING AND CODING
Figure 1: Final CAD cast Assembly in SolidWorks
Figure 4: Fritzing Diagram of System
Figure 2: Isometric View of Cast in SolidWorks
Testing after assembly and wiring included:
LCD Display needed adjustment from potentiometer
Adjustment needed a long period to tighten
Loosening of the cast felt more jolted than the tightening procedure
Figure 3: Bottom-Left View of Cast in SolidWorks
Material & Design Background
Protective Layer constructed from Thermoplastic Polyurethane (TPU)
TPU filament was used to 3D print protective cast layer
Cast Design constructed using SolidWorks
The design of the cast needed to take into account the size of an elbow and wrist for the respective diameter dimensions of the model
Figure 7: Proposed CAD Model of Hand Section
Fabricate a third layer to protect and conceal the wiring and electronics.
Fabricate the hand section so the cast can provide proper fixation of the wrist.
Further testing on the long-term durability and functionality of the splint.
Further testing for optimal strain gauge placement
Improvement of motor adjustment time
Photograph of assembled Smart Arm Cast with wires exposed
Cast Dimensions:
Wrist diameter: 54mm
Elbow diameter: 97mm
Length of the cast: 6 in
Uniform Thickness: 6 mm
Lofted Slot Height (Elbow): 45mm
Lofted Slot Height (Wrist): 30mm
LCD Case Interior: 32mm
Servo Motor Case Interior: 18mm
Battery
Supplies 4.5 Volt Power to Wires
(2x) Strain Gauge
Attached at either side of the patient’s forearm; Detects the pressure the cast puts onto the patient’s arm
Arduino Nano R3
Computer controller calculating pressure and electrical signals to operate motor and pulse oximeter
Switch Push Button
Placed so that the patient can recalibrate the tightness of the cast during the coding
16x2 Liquid Crystal Display
Displays heart rate and tightening pressure to patient
Tower Pro SG92R Servo Motor
Tightens and loosens cast using nylon wire and a pulley wheel
MAX30102 Pulse Oximeter
/
Detects heart rate by sending and receiving light waves in the patient’s blood vessels
Table 1: TPU Material Properties
Thermoplastic Polyurethane Material Properties |  | 
Property | Value | Description
Density [g/cc] | 1.45 | Light-weight, Low density
Hardness [R] | 66.3 | Durable, Hard
Max Tensile Strength [Mpa] | 63.3 | Durable; Can withstand high force and pressure
Elastic Modulus [Gpa] | 2.58 | Relatively low elastic modulus, easily stretches, does not deform
Elongation at Yield [%] | 23.8 | Amount material can stretch from its original state without breaking
Elongation at Break [%] | 26.6 | Amount material can stretch till breaking
Flexural Modulus [Gpa] | 2.18 | Relatively flexible; Low resistance to bending force
Water Absorption [%] | 0.24% | Low water absorption
Figure 6: Arm Cast Placed on Achilles’ arm performing a tightening calibration to return to baseline tightness
Figure 5: Serial Monitor of Arduino Calculating Heat Rate (Step 3)